F. 7.03-03

Besterekov U., Goncharov V.

COLLECTION OF METHODICAL INSTRUCTIONS

"DETERMINATION OF QUANTITATIVE AND QUALITATIVE INDICATORS OF WATER"

Shymkent ,2014

MINISTRY OF EDUCATION AND SCIENCE OF THE REPUBLIC OF KAZAKHSTAN

M.Auezov South Kazakhstan State University

Laboratory "Water quality monitoring of water resources of

the Republic of Kazakhstan"

Besterekov U., Goncharov V.

COLLECTION OF METHODICAL INSTRUCTIONS

"DETERMINATION OF                                                                                 QUANTITATIVE AND QUALITATIVE INDICATORS OF WATER"

Shymkent , 2014

UDC 536.7 : 54 ( 0.75)

Compiled by Prof. Besterek B ; undergraduate Goncharov V .

Collection of methodical instructions "Determination of quantitative and qualitative indicators of water." Shymkent: South Kazakhstan State University SKSU , 2014 . – 57 p .

The collection of methodological instructions given measurement procedure qualitative and quantitative indicators of water in the field, determine the concentrations of heavy metals in natural waters and super pure water for laboratory use.

Collection of methodical instructions intended for students , undergraduates, doctoral students , researchers research institutes and industrial enterprises.

Reviewer:

Anarbaev AA - Professor SKSU . M.Auezov

Considered at the meeting of the Department " CTIS " ( protocol number 6 of "_____" in _____________ 2014. ) And recommended for consideration by the methodical commission Faculty of Chemical Technology ( Minutes № ______ of "_______" in ___________ 2014. ) .

Recommended for printing Educational -Methodical Council SKSU im.M. Auezova ( protocol number ________ from "________" in ____________ 2014. ) .

© South Kazakhstan State University SKSU M.Auezov, 2014

CONTENTS

Part 1. Measurement of water quality in field condition with use multiparameter scanner PCD 650

| Introduction ……………………………………………………………..... | 8 |
| --- | --- |
| 1.Purpose and specifications multiparameter PCD 650…………………... | 8 |
| 2. Preparing for use……………………………………………………….. | 10 |
| 3 . Measuring modes and procedures for measurement ………………. | 11 |
| 4. Temperature measurement. Setting the measurement. Measurement Procedure…………………………………………………………………. | 11 |
| 5 . Measuring the level of PH. Setting the measurement. Calibration parameters PH level ……………………………………………………… | 12 |
| 6. Measurement the redox potential…………………………………… | 16 |
| 7. Measurement of the concentration of ions. Setting the measurement. Measurement Procedure………………………………………………….. | 16 |
| 8. Measurement conductivity of the solution. Setting the measurement. Measurement Procedure………………………………………………….. | 17 |
| 9. Measurement TDS of the solution. Setting the measurement. Measurement Procedure………………………………………………….. | 18 |
| 10 . Measuring the salinity of the solution. Setting the measurement. Measurement Procedure …………………………………………………. | 19 |
| 11. Measurement resistivity. Setting the measurement. Measurement Procedure ………………………………………………………………… | 20 |
| 12. Measuring the degree of saturation of the solution with oxygen. Setting the measurement. Measurement Procedure …………………….. | 21 |
| 13. Measurement of the concentration of dissolved oxygen in the water . Setting the measurement………………………………………………… | 22 |
| 14. Configure the device………………………………………………….. | 24 |
| 15. Error monitoring of measurement results……………………………. | 24 |

Part 2. Determination concentration of heavy metals in water by atomic absorption analysis using a the spectrometer AAnalyst 800

| Introduction………………………………………………………………...... | 24 |
| --- | --- |
| 1.Brief theoretical information……………………………………………… | 25 |
| 2.Methods of atomic absorption spectroscopy……………………………… | 26 |
| 3 . Work technique…………………………………………………………... | 28 |
| 3.1 Sample preparation……………………………………………………… | 28 |
| 3.2 Preparation of solutions for calibration………………………………….. | 28 |
| 3.3 Calibration of the instrument assessment and stability of the calibration curve…………………………………………………………………………. | 29 |
| 3.4 The general scheme of the analytical process with atomic absorption analysis………………………………………………………………………. | 29 |
| 4. Description and specifications of an atomic absorption spectrometer AAnalyst 800………………………………………………………………... | 30 |
| 5. Atomic absorption analysis with flame atomization elements…………… | 32 |
| 5.1 Sequence of of operations while working with flame atomization elements……………………………………………………………………... | 32 |
| 5. 2 Serial analyzes with flame atomization elements………………………. | 35 |
| 6.Atomic absorption analysis with electrothermal atomization elements…... | 36 |
| 6. 1 The general algorithm for the operation of a device……………………. | 36 |
| 6. 2 Serial analyzes with electrothermal atomization elements…………….. | 40 |

Part 3. Determination of mercury and arsenic in natural waters by using

flow injection analysis system FIAS 100 and  spectrometer AAnalyst 800

| Introduction ………………………………………………………………… | 42 |
| --- | --- |
| 1. Brief theoretical information ……………………………………………... | 42 |
| 2 . Technique with a flow- injection system FIAS 100 ……………………. | 43 |
| 2.1 Priority operations in preparation FIAS 100 system to work…………… | 43 |
| 2.2 Priority operations to optimize the absorption signal…………………… | 45 |
| 2.2.1 Preparing to Optimize…………………………………………………. | 46 |
| 2.2.2 Optimization of the absorption signal…………………………………. | 46 |
| 2.3 Implementation of serial analysis………………………………………. | 47 |
| 2.4 Priority operations during system shutdown FIAS 100…………………. | 48 |
| 3 . Recommended analytical parameters in determining arsenic and mercury in the samples…………………………………………. | 48 |
| 3.1 Using Reagents …………………………………………………………. |  |
| 3.2 Recommended analytical determination of arsenic parameters………… | 49 |
| 3.3 Recommended analytical parameters for determination of mercury with NaBH4………………………………………………………………………... | 49 |

Part 4. Preparation of high purity water using reverse osmosis systems ARIUM-61316

| Introduction………………………………………………………………...... | 49 |
| --- | --- |
| 1.Brief theoretical information……………………………………………… | 50 |
| 2.Application and system technical characteristics ARIUM-61316……….. | 51 |
| 3. Inlet Feed Water Requirement……………………………………………. | 52 |
| 4. RO System Schematic…………………………………………………… | 53 |
| 5.Operation of the reverse osmosis ARIUM- 61316…………………………. | 53 |
| 6.Procedure for switching on and off of the reverse osmosis………………………………………………………………………. | 53 |
| Requirements for qualification …………………………………………… | 54 |
| Safety regulations …………………………………………………………... | 54 |
| List of references…………………………………………………………….. | 54 |

Part 1. Measurement of water quality in field condition with use multiparameter scanner PCD 650

Introduction

In assessing the quality of water from natural surface water sources must first be on the ground water sampling is required to perform preliminary analyzes , since parameters such as temperature, pH, dissolved oxygen , salinity and conductivity, electrical resistivity as storage of water samples undergo significant changes.

Portable mnogoparametrny kiberskaner PCD 650 eutech ensures that field measurements of dissolved oxygen , electrical conductivity , salinity , resistivity , TDS (TDS), pH , redox potential , ion concentration and temperature.

This measurement procedure applies to all types of water sources.

Objective methods: determination of the content of dissolved oxygen, specific conductivity, salinity, resistivity, TDS (TDS), pH, redox potential, ion concentration and temperature on ground water sampling.

1.Purpose and specifications multiparameter PCD 650

Portable multiparameter scanner  PCD 650 eutech  provides exceptional flexibility for field measurements in various combinations of dissolved oxygen, electrical conductivity, salinity, resistivity, TDS (TDS), pH, redox potential, ion concentration and temperature. Also it - very powerful laboratory tool for measuring the above parameters

.

| Product Specification | Description |
| --- | --- |
| pH Range | -2.000 to 20.000 pH |
| Resolution & Accuracy | 0.1/0.01/0.001 pH & ±0.002 pH |
| No. of pH Calibration Points | 1 (Offset) to 6 points |
| pH Buffer Options | USA, NIST, DIN, PWB, Custom |
| mV Range | ±2000.0 mV |
| Resolution & Accuracy | 0.1 mV & ±0.2 mV |
| Ion Concentration Range | 0.001 to 19900 |
| Resolution & Accuracy | 2 or 3 digits & ±0.5% FS (mono); 1% FS (div) |
| No. of Ion Calibration Pts | Up to 8 |
| DO Range | 0 to 90.00 mg/L or ppm / 0 to 600% |
| Resolution & Accuracy | 0.01 mg/L or ppm / 0.1% ; ±0.2mg/L / ±0.2% |
| Salinity Correction | Automatic |
| Resolution & Accuracy | 0.0 to 50.0ppt & 0.1 ppt |
| Barometric Pressure Correction | Automatic correction with built-in sensor |
| Resolution & Accuracy | 450 to 825 mmHg / 59.9 to 109.9 kP & 0.1 mmHg or 0.1 kPa |
| Conductivity Range | .. to 500 mS/cm |
| Resolution & Accuracy | 0.05% FS & ±1% FS + 1 LSD |
| Cell Constant | 0.010 to 10.000 |
| Temperature Coefficient | Linear & Pure |
| TDS Range | .. to 500.0 ppt |
| Resolution & Accuracy | 0.05% FS & ±1% FS +1 LSD |
| Salinity Range | 0 to 80 ppt |
| Resolution & Accuracy | 0.05% FS & ±1% FS +1 LSD |
| Temperature Range | -10.0 to 110.0 °C / 14.0 to 230.0 °F (Meter) |
| Resolution & Accuracy | 0.1 °C/°F & ±0.5°C/ ±0.9°F |
| Temperature Compensation | Automatic / Manual (0 to 100 °C) |
| Power | 4 X 1.5 V "AA" Alkaline batteries or 9 V DC adapter, 500 mA |

2. Preparing for use

Attach the probes with correct type of connectors as indicated

Use the electrode properly for best results:

1) Keep the protective plastic electrode guard in tact during measurement and calibration. DO NOT REMOVE IT.

2) Always immerse the electrode beyond upper steel band as shown. Do NOT submerge the  electrode above the protective yellow cap.

3) Be sure to remove the protective electrode storage bottle or rubber cap of the pH electrode before calibration or measurement.

NOTE: If the electrode has been stored dry, wet the electrode in clean water for 10 minutes before calibrating or taking readings to saturate the pH electrode surface and to minimize drift.

3 . Measuring modes and procedures for measurement

Portable mnogoparametrny kiberskaner PCD 650 has the following measurement modes:

- Mode temperature measurement ;

- Mode measurement of pH ;

- Mode measurements of the redox potential .

- Mode of measuring the concentration of ions ;

- Mode measurements of the specific conductivity of the solution ;

- Measurement mode TDS;

- Mode measurement of mineralization of the solution ;

- Mode resistivity measurement solution;

- Measurement mode saturation of the solution with oxygen (O2%);

Mode - concentration measurement of dissolved oxygen in solution (O2 mg / l).

Measurement Procedure :

1.Prepare instrument for measurement.

2.Turn device by pressing F4. The device turns on for 3 seconds.

3.Pri turned the device will be in the measurement mode , which used to turn it off. Press F3 (MODE), to select the desired measurement mode ..

4.Pogruzit instrument probe into solution.

5.Pri completion of measurement , the display will light STABLE.

6.In the case if you want to keep the display measurement results need to press F1 (HOLD).

4. Temperature measurement. Setting the measurement. Measurement Procedure

The instrument is capable of measuring the temperature in the range of -10 to 110 ° C.

4.1 Setting the temperature measurement

To visit the temperature measurement mode settings, you must turn on the device, after press →, and then press F1 (SET), then ↑ and ↓ must choose a section TEM.

Page temperature measurement mode settings:

4.2 Temperature Measurement

1. Turn on the instrument by pressing F4. The device turns on for 3 seconds.

2 . Dip the tip of the device into the solution.

3 . Upon completion of the measurement , the display will light STABLE.

4 . Indicators to measure the temperature displayed on the screen in any measurement mode with all other key parameters .

5 . If you need to keep the display measurement results need to press F1 (HOLD).

5 . Measuring the level of PH. Setting the measurement. Calibration parameters PH level

Measurement Procedure .

The instrument is capable of measuring the pH value in the range from -2 to 20 pH units ;

5.1 Setting the measurement of pH.

To visit the measurement mode settings, PH level must include the unit , after press →, and then press F1 (SET), then press ↑ or ↓ to select the partition must RN .

Page mode setting measurement of pH

.

This page allows you to set the configuration of measuring the degree of acidity of the medium and the pH level calibration.

5.2 Calibration parameters PH level

Make sure you chose a standard buffer with which you want to calibrate the exponent acidity.

Calibration exponent acidity.

1. Turn on the device and make sure that it is in the measurement mode exponent acidity.

2. Pressing CAL (F2), to start the calibration process.

3.Pribor screen shows 'Calibration-Rinse Electrode' and then shows the calibration screen exponent acidity

.

4 . Rinse the probe in clean water and cook electrode calibration .

Note: If you want to completely re-calibrate the unit , you should clean

previous calibration data. Press CLR - C (F2), to clear the data from the previous

calibration. The instrument will display your screen confirmation. PRESS ENTER, to confirm the removal of the previous calibration data . Once data is cleared , you need to calibrate the first point , which corresponds to your chosen group buffer . Buffer group defaults - USA .

5 . Immerse the tube into the calibration buffer solution. The probe must be completely immersed in the solution .

6. The display shows the measurement of the exponent acidity. The instrument scans through all available ( uncalibrated ) point for the selected buffer group until it finds a close match within its calibration value . Once a match is found , the display shows the indicator 'Stable'.

7. Press ENTER, to confirm calibration.

8. Rinse the electrode with clean water.

9. Put the electrode into the second buffer solution and perform steps 5 and 7 to calibrate other points .

Calibration using custom solutions.

If the calibration of the instrument you are using custom solutions, then you should at least make 2 solution with the exact value of PH level .

1. Washed probe in clean water and cook electrode calibration .

Note: If you want to completely re-calibrate the unit , you should clean

previous calibration data. Press CLR - C (F2), to clear the previous calibration data . The instrument will display your screen confirmation. PRESS ENTER, to confirm the removal of the previous calibration data .

2 . Immerse the probe into the calibration buffer solution. The probe must be completely immersed in the solution .

3 . The screen shows two readings PH level . The top value shows the pH value from the previous calibration , and the lower the pH value indicates without any calibration. Use the ↑ and ↓, to adjust the display reading to riding pH level of your custom solution.

4 . Press ENTER, to confirm calibration.

5 . Rinse the electrode with clean water.

6. Put the electrode into the second buffer solution and perform steps 1 -4 to calibrate other points

Note: Function keys are available in gauge mode exponent acidity of the medium :

CLR-C (F2) - Clears the previous calibration data ;

NEXT (F3) - Displays the calibration report .

ESC (F4) - Exit the calibration mode and return to the measurement mode exponent acidity.

ENTER- Confirm the new calibration data .

5.3 Measurement Procedure :

1.Prepare instrument for measurement.

2.Remove the protective cap and rinse the electrode with clean water.

3.Vklyuchit device by pressing F4. The device turns on for 3 seconds.

4.Pri inclusion unit will be in the measurement mode , which used to turn it off. Press F3 (MODE), to select measurement PH.

5.Pogruzit instrument probe into solution.

6.When the completion of the measurement , the display will light STABLE.

7B if necessary to keep the display measurement results need to press F1 (HOLD).

6. Measurement of the redox potential

The instrument is capable of measuring the redox potential in the range from - 1999 to +1999 mV

Measurement Procedure

1. Prepare the device for measurement.

2 . Turn on the instrument by pressing F4.Pribor inclusive for 3 seconds .

3 . Press F3 (MODE), to select the measurement of the redox potential (mV).

4 . Dip the tip of the device into the solution.

5 . Upon completion of the measurement , the display will light STABLE

7. If you need to keep the display measurement results need to press F1 (HOLD).

7. Measurement of the concentration of ions. Setting the measurement. Measurement Procedure

Device capable of measuring the ion concentration in the range from 0 to 20,000 mg / l.

7.1. Setting the measurement.

To visit the measurement mode settings ion concentration must enable the device after press →, and then press F1 (SET), then press ↑ or ↓ to select the partition must ION.

Page measurement mode settings, the ion concentration.

This page allows you to set preferences measuring the concentration of ions.

7.2. Measurement Procedure

1.Prepare instrument for measurement.

2.Turn clicking F4.Pribor device turns on for 3 seconds.

3 . Press F3 (MODE), to select the measurement of the concentration of ions ION.

4 . Dip the tip of the device into the solution.

5 . Upon completion of the measurement , the display will light STABLE

7. If you need to keep the display measurement results need to press F1 (HOLD).

8. Measurement of specific conductivity of the solution. Setting the measurement. Measurement Procedure

Device capable of measuring the conductivity of the solution in the range from 0 to 200 mS / cm

8.1 Setting the Measurement Mode

To visit the measurement mode settings, the conductivity of the device you want to include , then press →, and then press F1 (SET), then press ↑ or ↓ to select the partition must Conductivity.

Page measurement mode settings, electrical conductivity of the solution.

This page allows you to set preferences measuring electrical conductivity of the solution.

8.2 . Measurement Procedure

1. Prepare the device for measurement.

2 . Turn on the instrument by pressing F4.Pribor inclusive for 3 seconds .

3 . Press F3 (MODE), to select conductivity measuring CONDUCTIVITY.

4 . Dip the tip of the device into the solution.

5 . Upon completion of the measurement , the display will light STABLE

7. If you need to keep the display measurement results need to press F1 (HOLD).

9. Measurement of salinity of the solution. Setting the measurement. Measurement Procedure

The instrument is capable of measuring the specific salt content of the solution in the range from 0 to 200 g / l.

9.1 Setting the Measurement Mode

To visit the measurement mode settings, you must include specific salinity device after press →, and then press F1 (SET), then press ↑ or ↓ to select the partition must TDS.

Page measurement mode settings, electrical conductivity of the solution.

This page allows you to set specific settings measuring salinity

9.2 . Measurement Procedure

1.Prepare instrument for measurement.

2.Turn clicking F4.Pribor device turns on for 3 seconds.

3 . Press F3 (MODE), to select the measurement of specific TDS TDS.

4 . Dip the tip of the device into the solution.

5 . Upon completion of the measurement , the display will light STABLE

7. If you need to keep the display measurement results need to press F1 (HOLD).

10 . Measuring the salinity of the solution. Setting the measurement. Measurement Procedure

Device capable of measuring the salinity of the solution in the range from 0 to 70 g/ l

10.1 Setting the measurement mode

To visit the measurement mode settings, mineralization of the solution must enable the device after press →, and then press F1 (SET), then press ↑ or ↓ to select the partition must Salinity.

Page measurement mode settings mineralization.

10.2. Measurement Procedure

1. Prepare the device for measurement.

2 . Turn on the instrument by pressing F4.Pribor inclusive for 3 seconds .

3 . Press F3 (MODE), to select the measurement solution salinity SALINITY.

4 . Dip the tip of the device into the solution.

5 . Upon completion of the measurement , the display will light STABLE

7. If you need to keep the display measurement results need to press F1 (HOLD).

11. Measurement of resistivity. Setting the measurement. Measurement Procedure

11.1 . Setting the resistivity of the of resistivity

To visit the measurement mode settings to the resistivity is necessary to turn on the device , after press →, and then press F1 (SET), then press ↑ or ↓ to select the partition must Resistivity.

Page mode setting resistivity measurements .

11.2. Measurement Procedure

1.Prepare instrument for measurement.

2.Turn clicking F4.Pribor device turns on for 3 seconds.

3 . Press F3 (MODE), to select a resistivity measurement solution RESISTIVITY.

4 . Dip the tip of the device into the solution.

5 . Upon completion of the measurement , the display will light STABLE

7. If you need to keep the display measurement results need to press F1 (HOLD).

12. Measuring the degree of saturation of the solution with oxygen. Setting the measurement. Measurement Procedure

The instrument is capable of measuring the degree of oxygen saturation of the solution in the range of 0 ... 500%.

12.1 Setting the measurement of oxygen saturation of the solution (O2%)

To visit the measurement mode settings ion concentration must enable the device after press →, and then press F1 (SET), then press ↑ or ↓ to select the partition you must DO saturation (%).

Page measurement mode settings, saturating the solution with oxygen.

This page allows you to set preferences measuring oxygen saturation of the solution.

12.2 . Measurement Procedure

1. Prepare the device for measurement.

2 . Turn on the instrument by pressing F4.Pribor inclusive for 3 seconds .

3 . Press F3 (MODE), to select the mode of measuring the degree of oxygen saturation of the solution DO (%).

4 . Dip the tip of the device into the solution.

5 . Upon completion of the measurement , the display will light STABLE

7. If you need to keep the display measurement results need to press F1 (HOLD).

13. Measurement of the concentration of dissolved oxygen in the water . Setting the measurement

Measurement Procedure .

The instrument is capable of measuring the concentration of dissolved oxygen in the solution in the range from 0 to 50 mg / l

13.2. Setting the measurement

To visit the measurement mode settings, the concentration of dissolved oxygen in the solution must enable the device after press →, and then press F1 (SET), then press ↑ or ↓ to select the partition you must DO Concentration (mg / L).

Page mode settings measuring the concentration of dissolved oxygen in solution (O2 mg / l).

This page allows you to set preferences measure the concentration of dissolved oxygen in the solution (O2 mg / l)

.

13.2. Measurement Procedure

1.Prepare instrument for measurement.

2.Turn clicking F4.Pribor device turns on for 3 seconds.

3.Nazhat press F3 (MODE), to select the concentration of dissolved oxygen in the solution DO (mg / l).

4 . Dip the tip of the device into the solution.

5 . Upon completion of the measurement , the display will light STABLE

7B if necessary to keep the display measurement results need to press F1 (HOLD).

14. Configure the device

Setting the value of the instrument is carried out only once during the first actuation of the device to work. Reconfiguration is performed in cases where the user needs to change the measurement depending on the characteristics of the object of research .

15. Error monitoring of measurement results

For best results, measure the pH level of the instrument must be calibrated in the following cases :

- When used regularly , at least once a week .

If not used , at least once a month.

- If you suspect that readings are incorrect.

- If corrosive liquids tested (very acidic or very alkaline) .

If liquid - tested in a wide range of measurement.

- After replacing the electrode.

Calibration is carried out PH level indicators according to the instructions described in Section 5.

Calibration measurements of other parameters is performed only if the replacement factory electrodes.

Part 2. Determination concentration of heavy metals in water by atomic absorption analysis using a the spectrometer AAnalyst 800

Introduction

In 1955, Australian scientist A. Walsh proposed a simple and practical way to easily workable quantitative determination of elements in solutions sprayed in the air -acetylene flame , by the absorption of radiation from the atomic lines special selective lamps.

It is now apparent simple solution, which is the basis of the analytical method of atomic absorption spectrometry determined the further rapid development method.

Atomic absorption analysis is pretty close to the traditional methods of "wet" chemistry, since the definition of the content of elements often revolves from solutions that in many cases provides preliminary chemical sample preparation . However , unlike most of the chemical methods , atomic absorption spectrometry

has a very high selectivity . Therefore, almost rarely required separation accompanying elements , as their presence does not usually cause noticeable systematic error definitions. This is due to the fact that the number of spectral absorption lines used is small, and they have a very small width of the order of 10-3-10-2 nm. As a result, the possibility of overlap of spectral lines of different elementsvery small. Therefore, the procedure for preparing samples for analysis is much simpler than conventional  methods for the "wet" chemistry, and rarely requires time-consuming analytical operations . The method allows to determine now about 70 elements. Basically this metal , but may use the method and some nonmetals : As, B, I, P, Se, Si , and Te. And, for the most determined elements ( 40 ) can be achieved relatively low detection limits : in a fiery version - usually from tenths to tens and hundreds of micrograms / liter , in electrothermal version - from thousandths to tenths of microgram / ​​liter.

1.Brief theoretical information

Atomic absorption spectroscopy (AAS) is a spectroanalytical procedure for the quantitative determination of chemical elements employing the absorption of optical radiation (light) by free atoms in the gaseous state.

In analytical chemistry the technique is used for determining the concentration of a particular element (the analyte) in a sample to be analyzed. AAS can be used to determine over 70 different elements in solution or directly in solid samples employed in pharmacology, biophysics and toxicology research.

Atomic absorption spectroscopy was first used as an analytical technique, and the underlying principles were established in the second half of the 19th century by Robert Wilhelm Bunsen and Gustav Robert Kirchhoff, both professors at the University of Heidelberg, Germany.

The modern form of AAS was largely developed during the 1950s by a team of Australian chemists. They were led by Sir Alan Walsh at the Commonwealth Scientific and Industrial Research Organisation (CSIRO), Division of Chemical Physics, in Melbourne, Australia.

The technique makes use of absorption spectrometry to assess the concentration of an analyte in a sample. It requires standards with known analyte content to establish the relation between the measured absorbance and the analyte concentration and relies therefore on the Beer-Lambert Law.

In short, the electrons of the atoms in the atomizer can be promoted to higher orbitals (excited state) for a short period of time (nanoseconds) by absorbing a defined quantity of energy (radiation of a given wavelength). This amount of energy, i.e., wavelength, is specific to a particular electron transition in a particular element. In general, each wavelength corresponds to only one element, and the width of an absorption line is only of the order of a few picometers (pm), which gives the technique its elemental selectivity. The radiation flux without a sample and with a sample in the atomizer is measured using a detector, and the ratio between the two values (the absorbance) is converted to analyte concentration or mass using the Beer-Lambert Law.

Atomic absorption spectrometer block diagram

2.Methods of atomic absorption spectroscopy

Flame Atomic Absorption Spectroscopy

Atomic Absorption (AA) occurs when a ground state atom absorbs energy in the form of light of a specific wavelength and is elevated to an excited state. The amount of light energy absorbed at this wavelength will increase as the number of atoms of the selected element in the light path increases. The relationship between the amount of light absorbed and the concentration of analytes present in known standards can be used to determine unknown sample concentrations by measuring the amount of light they absorb.

Performing atomic absorption spectroscopy requires a primary light source, an atom source, a monochromator to isolate the specific wavelength of light to be measured, a detector to measure the light accurately, electronics to process the data signal and a data display or reporting system to show the results. (See Figure 1.) The light source normally used is a hollow cathode lamp (HCL) or an electrodeless discharge lamp (EDL). In general, a different lamp is used for each element to be determined, although in some cases, a few  elements may be combined in a multi-element lamp. In the past, photomultiplier tubes have been used as the detector. However, in most modern instruments, solid-state detectors are now used. Flow Injection Mercury Systems (FIMS) are specialized, easy-to-operate atomic absorption spectrometers for the determination of mercury. These instruments use a high-performance single-beam optical system with a low-pressure mercury lampand solar-blind detector for maximum performance.

Whatever the system, the atom source used must produce free analyte atoms from the sample. The source of energy for free-atom production is heat, most

commonly in the form of an air/ acetylene or nitrous-oxide/acetylene flame. The sample is introduced as an aerosol into the flame by the sample-introduction system consisting of a nebulizer and spray chamber. The burner head is aligned so that the light beam passes through the flame, where the light is absorbed.

The major limitation of Flame AA is that the burner-nebulizer system is a relatively inefficient sampling device. Only a small fraction of the sample reaches the flame, and the atomized sample passes quickly through the light path. An improved sampling device would atomize the entire sample and retain the atomized sample in the light path for an extended period of time, enhancing the sensitivity of the technique. Which leads us to the next option - electrothermal vaporization using a graphite furnace.

Graphite Furnace Atomic Absorption Spectroscopy

With Graphite Furnace Atomic Absorption (GFAA), the sample is introduced directly into a graphite tube, which is then heated in a programmed series of steps to remove the solvent and major matrix components and to atomize the remaining sample.

All of the analyte is atomized, and the atoms are retained within the tube (and the light path, which passes through the tube) for an extended period of time. As a result, sensitivity and detection limits are significantly improved over Flame AA.

Graphite Furnace analysis times are longer than those for Flame sampling, and fewer elements can be determined using GFAA.

However, the enhanced sensitivity of GFAA, and its ability to analyze very small samples, significantly expands the capabilities of atomic absorption.

GFAA allows the determination of over 40 elements in microliter sample volumes with detection limits typically 100 to 1000 times better than those of Flame AA systems.

3 . Work technique

3.1 Sample preparation

Aqueous solutions can be introduced into the atomizer is often directly without any prior sample preparation. Water samples must be filtered and preserved . For the preservation of samples in the sample was added 0.005 mL of concentrated nitric acid in 1 mL of sample ( sample of 2.5 ml of nitric acid per 500 ml of sample ) .

3.2 Preparation of solutions for calibration

Procedure for the preparation of calibration solutions of CO

1. Prepare the necessary utensils .

- 7 flasks of 50 ml. glassware washed with conc. nitric acid, rinsed with tap water, rinsed with dist. water and dried in an oven at a temperature not lower than 1050C 2 . AAS prepare to work according to instructions.

3 . Prepare a solution with nitric acid concentration of 1 mol/dm3 (1H) . To this 63.7 cm3 conc . Nitric acid 1.413 g/cm3 density transferred into a volumetric flask with 1000 cm3 of distilled previously introduced . water ( 400-600 cm3) and bring to volume dist. water. The solution was stored in a glass container is not more than 1 year . Next , a solution concentration of 0.1 mol / dm 3 or 0.1 N.Dlya 100 cm3 of a solution concentration of 1 mol/dm3 (1H ) is transferred into a measuring flask of 1000 cm3 and the volume is brought up to the mark dist. water.

4 . Prepare standard solutions from basic GSO using solutions containing the metals solution 1 g/dm3 .

Originally prepared calibration solution A , the content of metal ions 100 mkg/sm3.To this volumetric flask of 50 cm3 of 5 cm3 GSO making metal solution and the volume is brought into the flask to the mark with 0.1 N HNO3. Then prepare the calibration solution in the content of metal ions 10 mkg/sm3 . To this volumetric flask with 50 cm3 of 5 cm3 make calibration solution A, and the volume in the flask was adjusted to the mark with 0.1 N HNO3. Solutions A and B were stored at ambient temperature in a closed form is not more than 1 year (ГОСТ 30178-96 "Raw materials and food products. Atomic absorption method for the determination of toxic elements ").

5 . Prepare a working calibration solution on the day of analysis in flasks with 50 cm3 in accordance with the data table below :

| № | Mass concentration of metals in the calibration solutions mkg/sm3 | Aliquot of solution, cm3 was placed in a volumetric flask with 50 cm3 |
| --- | --- | --- |
| 1 | 5,0 | A solution of 2.5 |
| 2 | 2,0 | A solution of 1,0 |
| 3 | 1,0 | В solution of 5,0 |
| 4 | 0,5 | В solution of 2,5 |
| 5 | 0,2 | В solution of 1,0 |

After introducing into the flask a solution of the metal in solution is adjusted volume flask to the mark with 0.1 N HNO3.

3.3 Calibration of the instrument assessment and stability of the calibration curve

Perform calibration device before starting measurements . For calibration curves repeatedly measured absorbance values ​​blank sample ( a solution of 0.1 N HNO3), and then each of the five calibration solutions sequentially for each element. Measurements were carried out in order of increasing concentration of elements in solution. The results of measurements automatically build the calibration curves according to the mean absorbance value of each element of its mass concentration of analytes in the calibration solutions. Stability control calibration curves carried out as follows : stability control calibration curve daily before measuring and during measurements is not more than 15 water samples analyzed . When this was analyzed 1-2 calibration solution. To control the stability of the calibration curve measured element concentration in the calibration solution , which lies in the middle of the measuring range of the element. The calibration graph is considered stable if the following conditions are:

(Сизм-Сзад)Сзад≤Δгр

where,, С изм – the measured value of the mass concentration in the same calibration solution, mkg/sm3;

С зад – set value  element concentration in the measured calibration solutions, mkg/sm3;

Δгр – total error of calibration, not more than 10%

If the calibration feature is not stable, find out and eliminate the causes of instability (not exactly prepared calibration solutions, not compliance with the conditions of preparation of the device in accordance with Table 3 .... techniques, etc.) and is repeated on the control element by using other samples for calibration. If you re not detected stability build a new calibration curve.

When changing reagents, long break in the instrument re-calibration of the instrument is carried out on all elements.

3.4 The general scheme of the analytical process with atomic absorption analysis

Atomic absorption method for analysis of the most developed to determine the elemental composition of liquid samples . On this basis, for the implementation of the analytical process perform the following steps .

1. Conduct representative sampling .

2 . Using headphones standard samples or reference samples , prepare a series of calibration solutions covering of the element needed for a range of the calibration curve .

3 . Prepared for operation atomic absorption spectrometer to acquire the signal in optimal absorption conditions of the investigated element .

4 . Enter the sample solution into the atomizer create absorbing layer atomic vapor and produce measurement of the analytical signal :

4 . 1. the solution is introduced into the atomizer , wherein the solvent evaporates, the dry residue is decomposed and converted to an atomic vapor ;

4 . 2 . through the atomic vapor layer transmit light from the light source selectively emitting a narrow resonance line characteristic of the element ;

4 . 3 . of the luminous flux transmitted through the absorbing layer , isolated part of the spectrum corresponding to the resonance absorption lines ;

4 . 4 . evaluate the analytical signal - absorption.

5 . Using calibration solutions prepared calibration dependence ( graphic or formulaic ) .

6. The concentration of the element under study in solution , and then in the sample.

7. Assess the validity of the definitions by comparing the results of analysis of standard samples of data for their certification .

4.Description and specifications of an atomic absorption spectrometer AAnalyst 800

Spectrometer AAnalyst 800 is a two-beam device , collected under the scheme Littrow with automatic wavelength selection and scanning. Spectral range 190-870 nm, the instrument is the nominal linear dispersion - 1.6 nm / mm . The device has three fixed slit width (0.2 , 0.7 and 2.0 nm) , implemented automatic selection of the height and width of the slit . In a fiery version uses a two-pronged scheme with deuterium background corrector . When using electrothermal atomization correction nonselective absorption by a longitudinal Zeeman effect with modulated magnetic field is applied single beam optical scheme . The detector of optical radiation in the instrument used wideband segmented semiconductor detector integrated with a low noise charge amplifier array . Appearance atomic absorption spectrometer AAnalyst 800 included in

electrothermal atomization form elements with an automatic dispenser samples

The device has an 8- tube holder with integrated power supply for the hollow cathode lamps and high-frequency electrodeless lamps. Control selection of lamps and their alignment relative to the optical axis of the device by a computer . In this recognition element radiation lamps and installation of analytical instrument parameters for

definition of this element automatically programmatically

instrument software .

Optical System

When working with flame atomization produced automated cost control of fuel gas and oxidant , as well as maintaining their ratio constant . Cycles ignition and extinguishing the flame acetylene dinitrooksid are automatic. Optimization of the vertical and horizontal position of the burner by a computer .

Interlocks prevent ignition of the burner , if not installed and spray with no drainage or insufficient pressure gas flame. Fuel supply is cut off the gas if the flame goes out or if any active lock.

When electrothermal atomization using graphite furnace heating cross (THGA). Provided use of graphite furnace with an integrated platform of Lviv, which improves the sensitivity and accuracy of the analysis , simplifies and accelerates the development of methods of analysis .

Software instrument is adjusted position graphite furnace to the optical axis , the speed of external and internal furnace blower protective gas. Temperature-time program heating furnace mozhetsoderzhat to 12 stages. Each stage can be individually programmed temperature to 2600 ° C , ramping to temperature for 0-99 (step 1 ) , the holding time at a constant temperature of 0-99 (step 1 ) , the gas flow through the

inside the oven cavity (0 . of 50 , 250 ml / min) with the possibility of using an alternative gas , programmed pneumatic opening and closing of the furnace. Provides a closed water cooling system of the furnace. Replacing the flame and electrothermal atomizers automatically. Autosampler contains tripod 88 or 146 positions for samples and comparison of solutions , as well as the flow capacity of the dosing system for flushing . The minimum sample volume for analysis is 0.1 ml. Dosed into a graphite furnace sample volume is 1-99 l (step 1mkl ) . The maximum dosing volume introduced into the furnace solution is 99 ml. Dosing system for flushing consumes 1.3 ml of solvent . Control of the automatic dispenser fully implemented by the instrument software .

In the spectrometer covers the full computer control of the instrument and sample introduction systems using a software package AA Winlab, collecting data processing and storage . The device operates with the absorption of -0.5 to 2.0, the concentrations of the elements and the intensities of the spectral lines. Signal integration time is 0.1-60 units of 0.1 s . perhaps

averaging signals , measurement of the area and peak amplitude signals , statistical processing of measurement results. In the quantitative analysis can be used up to 15 calibration solutions and conduct instrument recalibration one solution .

5. Atomic absorption analysis with flame atomization elements.

5. 1. Sequence of of operations while working with flame atomization elements

1. Turn exhaust ventilation.

2 . Turn on the air compressor and the supply of acetylene.

3 . Turn on the computer .

4 . Turn on the atomic absorption spectrometer . Before turning the spectrometer it must be installed at least one hollow cathode lamp (HCL). To do this:

4.1. open the lamp compartment by lifting the front and turning back ;

4.2 insert the lamp with a hollow cathode type Lumina for the desired item is selected in any holder and ensure that the plug is integrated fully into the lamp socket lamp holder ;

4.3 Close the lamp compartment ;

4.4 include atomic absorption spectrometer using the On / Off switch on the front panel .

5 . If there is a situation requiring emergency stop device , then quickly perform the following procedures :

5.1 turn spectrometer operating switch On / Off on the front panel ;

5.2 turn off gases from sources spectrometer .

6. Start WinLab32. When opening WInLab32 for AA automatically set stored earlier in the latter method, technique of atomic absorption measurements : flame (Flame), electrothermal (Furnace) or mercury- hydride (MHS) - this indicates the top line of the dialog box WInLab32. If necessary, replace the following techniques :

6.1ubedites that included an air compressor providing the necessary pressure for the tires;

6.2 File menu, click Change Technique, and then select the appropriate technique

7. Install necessary for hollow cathode lamps or high-frequency electrodeless lamp bulb holders free . Hollow cathode lamps may be connected in any of the holders 8 , and the high frequency electrodeless lamp holders only in the first to fourth . To install high-frequency electrodeless lamps:

7.1 open the lamp compartment by lifting the front and turning back ;

if necessary, assemble the lamp ( Figure) : insert generator unit high-frequency electrodeless lamps (driver) into the sleeve lamp selected item.

7.2 gently insert the lamp holder in the selected ( Only in the position of the first to fourth !) So that it rested against the back wall of the holder ;

7.3 Insert the plug corresponding to the selected high electrodeless lamp , into the outlet on the lamp holder ;plug the power cable into the appropriate socket bulb for high-frequency electrodeless lamps left of the lamp compartment.Use a socket with the same number as that of the holder ;

7.4 Close the lamp compartment.

8. To light the lamp to warm up , open with the icon on the screen Lamps settings window Align Lamps lamps and put the necessary lamp button On / Off. After ignition, the lamp appears green horizontal columnar histogram

light intensity on a scale AA. If the length of the histogram does not correspond to the middle of the scale , then click Set Midscale. Close the window. Give warm lamps before analysis work at least 10 minutes.

9. Open the computer screen method for user -defined element Me : File -> Open -> Method and select the appropriate method (OK). The display will show the name of the method, and defines the element .

14. Light the flame. To do this, click the icon on the screen Flame after Open Flame Control.

Ignite the flame in the window by clicking on the Flame Control Off / On.

15. If there was a change on the measurement technique sFurnace Flame, it is necessary to optimize the position of the torch .

15.1. Flame Control window , click Align Burner. This will open a window with a continuous signal recording Continuous Graphics graphics and master burner alignment Align Burner Wizard. Wait until the device is configured to check the absorption signal chosen element ( Continuous Graphics window appears continuous yellow signal from the selective light source ) , apply in a spray bottle with distilled water and press zero background signal Auto Zero Graph.

Select Align Burner Wizard window automatic settings optimal position burner Automatically align the burner , and then click Next. In the window that opens Determine Optimum Height press the power button operation Adjust. Watch the screen Continuous Graphics, as the device finds the optimal position and height of the torch at the end of the operation (Optimum height found - found the optimum height ) Click OK and Next.

15.2. Apply the solution to the atomizer element that you want to define a known concentration , the corresponding check concentration of the element , ie, giving optimal conditions the absorption signal is 0.2 (see Tools -> Recommended Conditions - meaning Sensitivity Check). To establish the optimum horizontal position of the burner in the window Determine Horizontal Position , press the power operation Adjust. Watch the screen Continuous Graphics, as the device finds the optimal horizontal position of the burner and after operation (Horizontal reference position found- standard position horizontal position found) Click OK and Finish.

16. Check the value of the characteristic element concentration . To do this, click on the menu Analyses Characteristic Concentration. Using the keyboard, enter into the corresponding line of the window values ​​used concentration of the element in a solution (Sample Concentration) and its maximum stored value of analytical signal (Instruments Readings). Click on a computer keyboard key Tab, and in the Characteristic Concentration measured value appears characteristic element concentration -Measured Characteristic Concentration. This value must be within 20% of the passport ( displayed on the screen - Comparison Characteristic Concentration).

17. If this condition is satisfied - you can run your serial analysis . If the value of the characteristic concentration significantly more passport , you must perform the following operations :

- Check the operation of the spray gun ;

- Check the preparation of the measured solution element;

- Work with a hollow cathode lamp ;

- Position adjustment of the burner ;

- Flow of gases.

18. Spend analysis in accordance with Section 5 . 2 .

19. After finishing the work , turn off the device .

19.1. Before switching off the device should be thoroughly flushed the entire system solution feeding the flames . Operations are carried out with a lit flame.

19.2 Flame Control window , click on the button Off / On.

19.3 Shut off the acetylene unit .

19.4 Release of acetylene residues device several times by clicking Bleed Gases in the Flame Control.

19.5Vyydite of the instrument software .

19.6 Turn the atomic absorption spectrometer .

19.7 Turn off the computer and printer.

19.8 Switch off the compressor .

19.9 Disconnect exhaust ventilation.

5. 2. Serial analyzes with flame atomization elements

Before starting the routine analysis should be performed all the operations turning the instrument conditions optimization analysis and verification of the characteristic element concentration ( Sec. 4 . 1). It is assumed that in memory already recorded previously developed methods for the user.

1. Users select the desired method to work . To do this, click the File menu Open -> Method. A window will appear Open Method. Select the desired contents on your method for a particular item and the type of products to be analyzed . Click OK.

2 . Click Wrkspc on display in the window that appears , select Workspace created and memorized earlier computer workstation for analysis ( absolutely necessary screens Manual Analysis Control, Results, Flame Control and Calibration Display). To create such a permanent computer workstation need to open these screens , pick their size and convenient place on the screen , but remember this setting : File Save As Workspace Name screen.

3.In the Manual Analysis Control window , specify the file name that will be written to all the analysis if necessary. For this row Results Data Set Name , click Open. A dialog box appears Select Results Data Sets. Select or enter the file name for the results of the analysis , and click OK . In the Manual Analysis Control in line During Analysis Save Data check box appears , indicating that all the results of the analysis will be recorded in this file.

4.Sozdayte calibration curve . Simply log in Method Editor  Calibration  Standart Concentration, then under Calib.Blank indicate acid ( write 0.5% HNO3). Then define standarty ( n / r : 0.2 0.5 1.0 2.0 5.0) in the graph Conc.

5.Podayte in sprayer control (background ) solution for the reference samples and Manual Analysis Control window , click Analyze Blank ( in full compliance with the information pointer on the right side of this button ) .

6.B Analyze Standard menu , select the first calibration solution for measurement. The program will introduce its concentration according imposed (given) method.

7.Podayte sprayer in the first calibration solution and click Analyze Standard for absorbance measurement.

8. § Repeat step 7 for all the calibration solutions .

9.Pristupayte to the analysis of samples. For this initial post control (background ) sample solution into a spray bottle and click Analyze Blank screen Manual Analysis Control.

9.1 Apply the first test solution into a spray bottle and click Analyze Sample screen Manual.

9.2 Following the called number matched to the analysis of samples , analyze all the planned solutions .

10 . When performing a large series of measurements in 15-20 samples is necessary to adjust the calibration curve . To adjust better to take the calibration solution of mid-range of the calibration curve and the solution (Reslop) to be specified in section Calibration Method Editor . Corrections during post graduation in flames this solution . In the entry field next to the button , select Analyze Standard name corresponding to the solution, and then click Analyze Standard. Also in 5-6 measurements necessary to analyze the background zperiodicheski solution samples Analyze Blank.

11. After measurements , perform the final operations off the device.

6. Atomic absorption analysis with electrothermal atomization elements

6. 1. The general algorithm for the operation of a device

1. Turn exhaust ventilation.

2 . Turn on the argon. Turn air compressor

3 . Turn on the computer .

3 . Turn on atomic absorption spectrophotometer ( On / Off button on the front panel ) .

4 . Start WinLab32. When opening WInLab32 for AA automatically set stored earlier in the latter method, technique of atomic absorption measurements : flame (Flame), electrothermal (Furnace) or mercury- hydride (MHS) - this indicates the top line of the dialog box WInLab32.

5 . When replacing a previously installed equipment flame (Flame) on electrothermal atomization , follow these steps :

5.1 Make sure that the air compressor provides the necessary pressure for the tires;

5.2 In the File menu, click Change Technique, and then Furnace. ON the instrument will change the techniques used atomization .

6. Software control window stove Furnace Control triggered by Furnace icon in the Tools menu and is used to control the furnace , air graphite furnace and adjusting automatic dispenser samples at a time when not carried out the analysis of samples . State of the furnace and the furnace program is shown in the top part of the window . Available in the window controls :

Furnace On / Off - starts or stops the furnace program , defined in the user;

Go to Next Step - stops the current program step furnace and starts the next step of the program;

Cleanout Temp. - Holding the cleaning oven at a selected temperature . For this operation, you need to dial the desired temperature furnace heating -and-click Start. For temperatures above 200 ° C the system maintains the temperature within 20 seconds, then turns off the heating . Heating furnace can be stopped manually, again by clicking the Start; Open / Close - opens or closes the oven with a pneumatic system;

Align Tip - displays the program Align Autosampler Tip, used to align the tip of the pipette automatic dispenser samples ; Condition Tube - runs a program consistent heating furnace at increasing temperatures for air furnace ; Flush Sampler - starts or stops a series of automatic washing steps autosampler . These operations must be performed by charge wash bottle or replace any pipe cleaning system ;

Align Furnace - runs a special program for automatic adjustment of the electrothermal atomizer relative to the probe beam selective light source.

7. Set the desired hollow cathode lamps or high-frequency electrodeless lamp in free slot bulb holder unit.

8.For turning the lamp to warm up , open with the icon on the toolbar Lampe corresponding window and turn on the lamp by pressing the desired On / Off. When this lamp will be energized nominal current passport . After ignition, the lamp appears green horizontal histogram columnar light intensity scale in AA. If the length of the histogram does not correspond to the middle of the scale , then click Set Midscale ( mid- scale installation ) . Give warm lamps before analysis at least 10 minutes.

9. Open method on a computer screen for user -defined element Me : File -> Open -> Method and select the appropriate method (OK). The display will show the name of the method, and defines the element .

Connecting the Drain Tube and Waste Bottle

10 . When changing equipment necessary to adjust the position . To adjust the position of the oven , follow these steps :

10.1 Click on the button in the Align Furnace Furnace Control. This opens the continuous graphic recording signal Continuous Graphics window and adjust the position of the master graphite furnace Align Furnace / Quartz Cell Wizard;

10.2 Click in the Continuous Graphics button auto zero Auto Zero Graph.

10.3 Select the window that appears Align Furnace / Quartz Cell Wizard mode automatic adjustment Automatically align the device and click Next.

11. In the opened window, click Adjust to determine the optimum geometrical position (Determine Position) electrothermal atomizer and Continuous Graphics window watch the device automatically finds the optimal vertical and horizontal position of the beam relative to the atomizer from the selective light source. After the operation , click Finish.

12. Configure the autosampler samples . You should check and adjust the alignment after moving automatic dispenser or changing any of its components, after replacing the graphite tube or contacts , or opening or closing the oven after adjustment graphite furnace .

The correct position of the metering capillary into a graphite furnace for introducing the sample solution on the furnace wall or platform

When properly configured, the tip of the pipette should take the right solution from a cup on the autosampler tray , and then inject the correct solution in the graphite tube . Usually the tip of the probe should be set as follows ( Fig. 5 . 4) that he stayed at a height of 2 mm above the surface of the wall or the platform of the city. If you use a very small sample volumes , it is best to adjust the tip of the probe so that it stayed at a height of only 1.0-1.5 mm above the wall or platform . The tip of the probe in any case must not hurt the furnace wall or platform . To configure automatic dispenser of samples ( adjust the horizontal position and depth of the pipette tip in the graphite tube and cup with the sample solution ) , follow these steps :

12.1. Be sure to set the system flushing. From the Tools menu , click Options Furnace Autosampler, select the size of the sampler tray , which you intend to use, then click OK.

12.2 Click in the Furnace Control button Align Tip. Wizard appears alignment automatic dispenser tip samples Align Autosampler Tip Wizard. Use it to specify further adjust the position of the tip of the sampling capillary samples in automatic dispenser cups solutions and dispensing opening of the furnace .

13. Flush the system . In the window, click on Furnace Control Flush Sampler. This activates a special washing cycle .

Repeat this cycle as many times until all of the tube will not be completely filled with rinsing liquid . Make sure that the system flushing are no air bubbles .

14. Allow the lamp to warm up for at least 10 minutes.

15. Fill the tray method editor solutions .

From the Tools menu , click on the Auto button and Results. Configure Automated Analysis Control window right method name and position in the sampler tray standard solution : Sample Info File Locations.

16. In the Automated Analysis Control page, click Analyze.

17.Provedite analysis in accordance with Sec. 5 . 2 .

18. To turn off the appliance , follow these steps :

18.1 Thoroughly rinse the entire supply system solutions in a graphite furnace . In the window, Furnace Control click Flush Sampler;

18.2 exit the software of the device;

18.3 turn atomic absorption spectrophotometer ;

18.4 Turn off the computer and printer ;

18.5 Turn off the compressor ;

18.6 Close the bottle with argon .

6. 2 . Serial analyzes with electrothermal atomization elements

1. Before the serial analysis should be performed all switching device operation , analysis, and optimization of the conditions of the characteristic verification element concentration . It is assumed that in memory already recorded previously developed methods for the user.

2 . Users select the desired method to work . To do this, click the File menu Open -> Method. A window will appear Open Method. Select the desired contents on your method for a particular item and the type of products to be analyzed . Click OK.

3 . Click Wrkspc on display in the window that appears , select Workspace created and memorized earlier computer workstation for analysis ( absolutely necessary screens Automated Analysis Control, Results, Transient Peaks Display and Calibration Display). To create such a permanent computer workstation need these screens , pick their size and convenient place on the screen and remember this setting : File Save As Workspace Name workplace.

4 . Create an information file about the samples :

4.1. In the File menu, click New -> Sample Info File.

4.2. In the window that opens New Sample Information File click on

button OK. Editor opens the sample data files

(Sample information Editor). Fill in the information file the necessary information about a series of samples that must be analyzed . Units for use w / w need to ask :

- A unit volume (Volume Units);

- Unit weight (Weight Units);

- The identifier ID ( identification) and sample mass samples (Initial Sample Wt.);

- Dilution of samples : the volume of the first solution by dissolving the sample (Sample Prep. Vol.), The volume of aliquot taken from him (Aliquot Volumes) and the final volume of the sample solution (Diluted To Vol.).

4.3. Remember the file under any name File -> Save As -> Sample info File.

5 . Tap in the Automated Analysis Control button Sample Information File: Open ... In the window that opens , select Open Sample Information Required information file and click Open.

6. In the Automated Analysis Control , specify the file name that will be written to all the analysis if you want to keep them . For this row Results Data Set Name , click Open. A dialog box appears Select Results Data Sets. Select or enter the file name for the results of the analysis , and click OK . In the Manual Analysis Control in line During Analysis Save Data check box appears , indicating that all the results of the analysis will be recorded in this file.

7. Load tray automatic dispenser samples. Make sure that you have the solutions to the provisions specified in the method editor and file information about samples. If you have a sample solutions with the addition of elements determined , then put them in the position immediately following the initial sample solution without additives and in the same order in which they are defined on page Calibration Method Editor .

8. Create a calibration curve . To do this, press the button Method Editor, from there choose Standard Concentrations, then Calculate Standard Volumes, then in Stock Standards specify the cell number and volume of which will prepare solutions of a predetermined concentration . In Location Blank indicates the number of cells , then specify standards . Standard Concentrations will come in all concentrations . In the Auto Sampler specify the cell where the diluent (Diluent Location 124)

8.1 in the Automated Analysis Control click Analyze ( at the bottom ) and in the next screen click on Calibrat. The instrument will measure the reference (blank ) solution for the calibration solutions , all calibration solutions and create a calibration curve , which can be observed in the Calibration Display.

9. If you are satisfied with the calibration , then proceed to the analysis of samples. To do this, Automated Analysis Control window , click Analyze Samples. The system will analyze all solutions as defined in the data file samples.

10. When using the method of standard additions in the Automated Analysis Control click Analyze ( at the bottom ) and in the next screen , click Analyze All. The instrument will measure the control (blank ) of a solution of the sample solution without the addition of the element , all the sample solutions containing additives of the element and create a calibration curve method, additives which can be observed in the Calibration Display, and determine the content of the element in the original sample . Then these operations are repeated for all sample solutions that you have chosen in the Automated Analysis Control: SetUp.

11. When performing a large series of measurements in 15-20 samples is necessary to adjust the calibration curve using the solution to control the position of the calibration curve (Reslope Std.). To adjust better to take the calibration solution of mid-range of the calibration curve . For automatic analysis on pages Calibration Method Editor you must determine the type and frequency of inspection and grading tilting calibration curve.

12. After measurements , perform the final operations off the device.

Part 3. Determination of mercury and arsenic in natural waters by using

flow injection analysis system FIAS 100 and  spectrometer AAnalyst 800

Introduction

In addressing important applications of modern analytical chemistry is often necessary to perform a large number of individual determinations . This is typical, for example , when determining the polluting substances in the environment , in clinical laboratories when performing biochemical investigations. A big role to play in this automatic methods of chemical analysis , using modern ideology and analytical tools and microprocessor technology . Development of automatic methods for analyzing stimulated increase reliability and efficiency analysis while reducing its cost, promptly obtain a large amount of accurate analytical information. Since the data necessary to process a lot, convenient to use proper microprocessor technology - in modern devices is part of a computer device . Such devices shall be suitable for automatic execution critical stages of analytical determination of sampling, transfer it into a form suitable for analysis - a solution for analytical reactions , measuring the amount of the product of this reaction .

There are two approaches to the automation of analytical determinations in solution. Realizing first seek to create an automated analyzer that would fully or partially fulfilled the usual analytic work. It is difficult and expensive. The simpler way and construct the device automatically performs only one important analytical operation - the actual measurement , such as titration. Appropriate devices called titrators .

More successful was the second , a fundamentally different approach , based on the idea of continuous analysis. Here, the sample solution is moved continuously with a peristaltic pump for fluid communication instrument - often just a system is glass or plastic tubes. When this is added to it the necessary reagents , and the reaction itself proceeds in the analytical stream. Determination of the concentration analysis of the reaction product - the analytical signal - is carried out continuously by means of the detector when the portion of the reaction product stream passes it. The signal recorded with the device as a peak . In this case, conditions are achieved most simply ensure reproducibility of measurements and analytical reactions . Method proved to be a convenient , simple and effective.

- Brief theoretical information

FIAS systems include one or two independently controlled flows of the peristaltic pump for a carrier of the sample and reagents , the capacity for chemical reactions , gas / liquid separator , a quartz cell ( for flame AA) . The sample is pumped through the metering pump and fed into a loop channel , which is constantly flushed with a solution of a carrier (typically , a dilute acid solution ) . For volatile compounds or elements gidridobrazuyuschih mercury vapor discharge using NaBH4 in the reducing solution or SnCl2, which are mixed with the sample in the container and the separator separates gas from the sample matrix hydride . Then argon flow analyzed portion of the sample is fed into the analysis unit .

The technique is based on the fact that when located in the unexcited ( ground) state absorbs photon energy specific ( intrinsic ) , and the energy in an excited state. The amount of energy of the characteristic radiation absorbed by a substance increases with the number of atoms of the element in the path of the light beam . The relationship between the amount of absorbed radiation and the element concentrations in the standards of a known concentration can be used to determine concentrations in unknown samples by measuring the radiation absorbed by them .

To perform atomic absorption analysis requires primary source

characteristic radiation source atomic spectrometer- monochromator , releasing light of a certain wavelength detector accurately detect radiation , electronics and data processing terminal or reporting system for reporting results . As the source of the characteristic used highly sensitive single beam optical scheme with low pressure mercury lamp . For each element uses a separate lamp.

2 . Technique with a flow- injection system FIAS 100

2.1 Priority operations in preparation FIAS 100 system to work

1. Turn exhaust ventilation.

2 . Turn on the air compressor .

3 . Turn on the acetylene and argon.

4 . Start your computer and WINDOWS.

5 . Turn on the atomic absorption spectrometer . before switching

spectrometer it must be installed at least one hollow cathode lamp (HCL).

6. Start WinLab32. To do this on the desktop , click the WINDOWS :

Start -> Program -> PerkinElmer WinLab32 for AA -> WinLab32 for AA or click on the icon WinLab32 for AA on your computer screen . On the display dialog box will appear with the software title at the top of the screen : WinLab32 AA Flame.

When opening WInLab32 for AA automatically set stored earlier in the latter method, technique of atomic absorption measurements : flame (Flame), electrothermal (Furnace) or flow- injection (Fias) - this indicates the top line of the dialog box WInLab32.

7. Select the method of FIAS.

8. Open all the windows you want to work . File → Open → Work space → 123 ( file ) → open .

9. Turn right lamp . To light the lamp to warm up , open with the icon on the screen Lamps settings window Align Lamps lamps and put the necessary lamp button On / Off. When this lamp will be energized nominal current passport and On / Off button will be green . After ignition, the lamp appears green horizontal histogram columnar light intensity scale in AA. If the length of the histogram does not correspond to the middle of the scale , then click Set Midscale ( mid- scale installation ) . While appropriately change the signal gain . Close the window. Give warm lamps before analysis at least 40 minutes .

Do not disconnect or remove the lamp when it is included . To turn off the lamp in the Align Lamps click on the On / Off, that the button has ceased to be green , which corresponds to powering off the lamp. Close the Align Lamps. Disconnect and remove the bulb . To remove the high-frequency electrodeless lamp as encoding shteker.6 .

10 . A quartz tube . Detection tests for arsenic , remove the side windows of the quartz tube . For analyzes of mercury side windows can not be removed .

11. Open method on a computer screen for user -defined element Me : File -> Open -> Method and select the appropriate method (OK). The display will show the name of the method, and defines the element .

12. Optimize the quartz tube . To do this, In the Flame Control click Align Burner. This will open a window with a continuous signal recording Continuous Graphics graphics and master burner alignment Align Burner Wizard. Wait until the device is configured to check the absorption signal chosen element ( Continuous Graphics window appears continuous yellow signal from the selective light source ) . Select Align Burner Wizard window automatic settings optimal position burner Automatically align the burner , and then click Next. In the window that opens Determine Optimum Height press the power button operation Adjust. Watch the screen Continuous Graphics, as the device finds the optimal position and height of the torch at the end of the operation (Optimum height found - and found the optimum height Horizontal reference position found- standard position horizontal position found) when finished click OK and Finish.

13. Be sure to connect the tubes properly made

14. Disconnect the feed sample , connecting the gas-liquid separator and the quartz tube . Place the free end of the tube in a beaker to drain all the liquid enters the tube and leave it in a glass.

15. Set the value of the carrier gas flow in accordance with the present in the " Recommended Operating Conditions " . Use of the gas flow regulator and flow meter ( rotameter ) on the front panel . If the tide is out , can be closed automatic gas valve. To enable flow in the FIAS control -click on Valve Fill / Inject.

Turn on and set the carrier gas stream before turning pumps . This prevents fluid from entering the check valve.

16. Lower the input ends of the pipes feeding the media ( yellow / blue ) , reductant ( red / red ) and the sample solution ( leading to the valve FIAS) in containers with deionized water.

17.Povernite levers pump tube to squeeze the pump to the rollers .

18. In the FIAS Control -click on Valve Fill / Inject, to set the valve to the Filling (Fill).

19. Enter a value for the speed of 100 Pump # 1 Speed

20. Turn on the pump by clicking on Pump # 1 \

21. Make sure there are no leaks in the system .

22. Set the flow of the carrier and reducing agent so that the ratio of the carrier to the flow of reducing agent was approximately 2:1 with a carrier fluid flow from 9 to 11 ml / min. To do this, follow the steps 23-30 .

23. Fill the graduated cylinder deionized water.

24. Immerse the inlet end of the tube feeding the media in a graduated cylinder .

25. After 1 minute, check volume reduction . The flow should be between 9 to 11 ml / min.

26. If the flow of the carrier beyond these limits , adjust the pressure on the pump tube until the value of the flux does not fall within the specified interval .

27. Fill the graduated cylinder deionized water.

28. Immerse the inlet end of the feed tube of a reducing agent in a graduated cylinder .

29. After 1 minute, check volume reduction . The flow should be approximately equal to half the carrier flow - between 5 and 7 ml / min.

30. If the flow of reductant beyond these limits , adjust the pressure on the pump tube until the value of the flow is in the desired range .

31. Adjust the drain of the gas-liquid separator. To do this, follow the steps 32-36

32. Place the reductant inlet ends of the tubes and the carrier into a vessel with deionized water.

33. Adjust the pressure on the drain pump tube to adjust the speed of flow of liquids from the gas- liquid separator . Increase pressure on the pump tube until the liquid flowing from the separator is not bubbles begin to appear .

34. The rate at which liquid flows out of the separator must be somewhat larger inflow rate of fluid into the separator.

35. Make sure that the supply tube sample is clean and dry . Connect the supply pipe sample from the separator to the quartz tube .

36. Turn off the pumps : FIAS Control window , click on the Pump # 1.

2.2 Priority operations to optimize the absorption signal .

Optimize signal each time when starting new analyzes and when you change the configuration of the liquid system .

2.2.1 Preparing to Optimize

1. Set up equipment , install and adjust the flow tube as described in the previous section . Make sure that the supply tube is connected to the sample quartz tube .

2 . Prepare a "blank " solution recommended by the carrier and reductant solutions and recommended solution is to test the sensitivity . See box recommended conditions ( from the Tools menu select Recommended Conditions) or section .

3 . In the Toolbar , click on MethEd. Open or standard method , or a method to create the user.

4 . Use the Method Editor to set the following parameter values ​​:

- On the set value Inst Read Delay 0 seconds.

- Inst page to adjust the Read Time 15 seconds.

- Page 5 for Calib enter Replicates.

7. Place the inlet end of the tube supplying the solution carrier ( yellow / blue) in the carrier solution , and the input end of the reductant feed tube (red / red), to a solution of a reducing agent.

8. Place the inlet end of the feed tube in the sample "blank " solution .

9. In the Manual Analysis Control -click on the Analyze Blank.

10 . Place after the input end of the tube feeding the sample solution in comparison.

11. In the Manual Analysis Control -click on the Analyse Sample or FIAS Control window , click on the FIAS On / Off. The signal from the first replica will appear in the Display Peaks.

12. Canculate intrinsic mass. To do this, follow the steps

13. Analyses from the menu , select Calculate Charact. Mass ... ( calculation of the characteristic mass). A window calculation of the intrinsic mass.

14. Enter the data for the standard and "blank " box computing solutions in the characteristic mass.

Data on standard and " idle " solutions are arranged in the data representation.

15.Click the Tab key after entering values.

The system calculates and presents the characteristic mass on the screen . The measured value must be no more than 20% higher than the values ​​given above . If it is very high, you need to optimize assay conditions

2.2.2 Optimizing signal swallowing

Maximum peak appears after 5-8 seconds after the Read Time and background signal must reach before the end of the line Read Time.

- If necessary, the adjustment may be performed as described below. After each setting wait for the next signal before carrying out further customize .

- Change in each case only one parameter.

If all replicas are made before you have finished setting , click the Analyze Sample or FIAS On / Off, to run the program again . Repeat this as many times as necessary.

- After calculate intrinsic weight as described in section 2.1.

The optimization procedure signal

1. Absorbance values ​​of all replicas must be close . If the first replica signal absorption higher than the next , lengthen stage Fill FIAS- page method. If the value of the absorption of the first replica below lengthen stage Prefill.

2.Make sure that the values ​​of the parameters Read Delay ( 0 sec. ) And Read Time ( 15 sec.) Method on page Inst installed correctly.

3. Adjusting the carrier gas flow may improve sensitivity. If the maximum peak appears too early , slightly reduce the flow of carrier gas . If the maximum peak appears too late , slightly increase the flow of carrier gas .

Note: If the carrier gas flow is too great , the mercury vapors dissipate too rapidly. If the carrier gas flow is too low, the mercury vapor in the cell arrive too slowly. In both cases, the signal strength and the sensitivity is low. Usually sufficient flow of 40-70 ml / min.

4. Decrease in the flow of fluid from the gas- liquid separator can increase sensitivity.

Note: If the effluent stream from the separator is too large, mercury vapor can fly through the drain hole . If the stream is too small , the liquid level can rise and it gets into the feed tube of the sample and FIMS- cell. If liquid enters the cell , the cell must be cleaned as described in "FIMS: Installation, Operation , description of the system . "

5. Flow regulation carrier and reductant solutions may increase the sensitivity .

6. Calculate again the characteristic mass and proceed to the analysis.

2.3. Perform serial analysis

1. Prepare for work in accordance with section 2.1

2 . Prepare solutions of the reducing agent and the carrier , and sample solutions graudirovochnye solutions . Recommended solutions , see

3 . Create or open method .

4.Pomestite inlet end of the feed tube samples in a "blank " solution .

5.In the Manual Control window , click on the Analyze Blank.Dlya stop analysis , click again to Analyze Blank.

6. Check the location and magnitude of the concentrations of solutions that will be analyzed after pressing the Analyze Standart, against the values ​​you entered on page Calib in the Editor window method.

7.Pomestite feed tube inlet end of the sample in the first calibration solution .

8.Ubedites that the number and magnitude of the concentration of the standard solution (Analyze Standard) correspond to a solution that you will analyze . Click the Analyze Standard. To stop the system , again click on Analyze Standard.

8.Povtorite these steps for each calibration solution .

9.If you want to save the calibration methods: From the File menu select Save as  Method, click on OK .

10 . Place the inlet end of the sample feed tube into the sample solution .

11.Click on the Analyze Sample.Chtoby stop the system , click again on the Analyze Sample.

2.4 Priority operations during system shutdown FIAS 100

1. Rinse the sample introduction tube , and the reagent carrier with deionized water.

2 . Place the inlet end of the sample supply tube , the carrier and reactant ( reducing agent , a buffer ) in vessels with deionized water.

3 . In the Toolbar , click the FIAS, a window FIAS Control.

4 . In the FIAS Control:

- Click on Valve Fill / Inject to valve position stood Fill ( valve position shown on the display in the window Status);

- Pump 1 to set the value 120. Then click on Pump 1 to turn on the pump ;

5 . In the FIAS Control -click on Valve Fill / Inject several times during the operation of the pumps . This will ensure effective flushing flow channel samples and internal valve parts FIAS.

Rinse the tube to completely remove all traces of the reagents used .

6. After washing with deionized water :

- Remove all the tubes of the vessels with washing solutions ;

- Let the pump run until all the tubes and gas-liquid separator is not freed from the liquid;

- Click on the Pump 1 to turn off the pump;

- Ensure that the valve is in the FIAS Fill ( Filling ) .

7. Unscrew all clamping levers holders of tubes.

8. Pour the spent liquors from the container to drain waste and remove traces of leaks. When disposing of solutions according to the internal rules of safety .

9. Shut off the carrier gas.

10 . Turn off the device :

- Exit AAWinLab;

Turn - FIMS- spectrometer ;

- Turn off the computer and printer.

11. Turn off the ventilation system.

3 . Suggested parameters for the analytical determination of arsenic and mercury in the samples

3.1 Using Reagents

Standard solutions :

Calibration solutions used with the systems FIAS-100/400 can be prepared by diluting the original factory ( head ) solutions. All solutions should be stored under an inert plastic container . Standard solutions with concentration elements 100 mg / l can normally be kept to a year. Solutions at concentrations below 1 mg / l must be prepared directly in the day of use . The stability of solutions of standard mercury largely influences on the adsorption walls used to store dishes. To stabilize the solution Hg recommended to use in their preparation mixture containing 2 percent by volume of nitric acid and sulfuric acid with the addition of 1-2 drops ( 30-60 ul) 5 % solution of KMnO4.

Carrier solution

For all elements analyzed , except for tin and mercury as a solution of the carrier is used a 10% solution of hydrochloric acid (10 parts of concentrated HCl at 90 parts water). When determining mercury recommended 3% (vol ) solution of hydrochloric acid. Carrier solution should always contain the same acid and about the same concentration as the test samples .

Recovery

As the reducing agent used is typically a freshly prepared aqueous solution containing 0.2% NaBH4 and 0.05% NaOH. Filtration is desirable to produce a reducing solution using vacuum filtration .

3.2 Recommended analytical parameters for determination of arsenic

Flow- injection system

1. Cell temperature : 900  C

2 . Pump speed : 100

reagents

Carrier: 10 % (by ​​volume ) HCl

The reducing agent : 0.2% NaBH4 in 0.05% NaOH

The sample solution : As3 + 10 % (by ​​volume ) HCl

Checking Sensitivity

At a dosage of 500 l of a solution with a concentration of arsenic 10 mg / l should be a signal at 0.45 A. The measurements were carried out using EDL. When using LPC signal about 0.30 A.

Notes

Increasing the concentration of the reducing agent ( about 0.5%) can lead to increased sensitivity.

Status of the quartz cell has a significant impact on the detection sensitivity . If the measurement sensitivity is low, the cell must be conditioned .

As analyzed by content samples should be subjected to the procedure pre-reduction . As the reducing agent solution is used , containing 5 % (w / v ) KI and 5% ( w / v) ascorbic acid. The analyzed samples or standard solutions before analysis should be maintained with a solution of a reducing agent at room temperature for 30-60 minutes.

The above-mentioned sensitivity meets the definition of As3 +. If arsenic is present in the form As5 +, observed sensitivity is 5-10 % higher than a given value . Therefore the samples must be restored prior to analysis to As3 +. Recovery rate increases with increasing acid concentration. Suitable for a successful recovery following recovery procedure .

To 1 ml sample or standard solution is added 1 mL of concentrated HCl and 1 mL of a solution containing 5 % (w / v ) KI and 5% ( w / v) ascorbic acid. The solution was kept for 45 minutes at room temperature and diluted to 10 ml.

If the test solution contains other oxidants may be necessary to add a larger amount of the reducing solution .

3.3 Recommended analytical parameters for determination of mercury with NaBH4

Flow- injection system

1. Cell temperature : 0  C

2 . Pump speed : 100

reagents

Medium: 3% (by ​​volume ) HCl

The reducing agent : 0.2% NaBH4 in 0.05% NaOH

The sample solution : Hg2 + in acidic medium

Checking Sensitivity

At a dosage of 500 l of a solution with a mercury concentration of 10 ug / l of the signal should be at 0.07 A.

Notes

To achieve the stated sensitivity of the carrier gas flow must be increased to 70-100 ml / min.

To stabilize the mercury in the analyzed samples and standard solutions to 100 ml should be added 1-2 drops of a 5% solution of KMnO4.

HCl concentration should be maintained at a minimum level to prevent recovery KMnO4. Usually suffice 1 ml concentrated HCl per 100 ml sample. When the mercury content of less than 10 mg / l can be observed sorption of mercury on the walls of containers used . Sorption is dependent on the material used must therefore be pre tableware verified the possibility of sorption.

Use of the reducing solution with a low concentration (e.g. 0.02% NaBH4 in 0.005% NaOH) can reduce signal interference and a reduction of the blank .

Part 4.  Preparation of high purity water using reverse osmosis systems ARIUM-61316

Introduction

Water of high purity are used every day in each laboratory in various applications , production and / or production of the reagents , the preparation of buffer solutions , nutrition of laboratory equipment , etc. . Depending on the required tasks at different water quality

Until recently, the most common method of obtaining water in the laboratory is the method of distillation. However, this method has several drawbacks . First, the use of distillation requires high operating costs : energy required for heating and vaporization of water is about 3000 kJ per kg steam , costs for water vapor condensation of about 8 liters of water 1 kg of steam . Second, the practical experience of leading experts shows that the water produced by this method is not always consistent with the requirements of apyrogenicity and physico- chemical parameters ( conductivity , dry residue , chlorides , ammonia , reducing substances , total organic carbon , etc.) . Maintenance costs obsolete distillation equipment often exceed the initial investment to replace them . As a result of all this , currently , instead of obsolete units distsilyatsionnyh modern water purification systems used for reverse osmosis.

1.Brief theoretical information

The process of osmosis through semipermeable membranes was first observed in 1748 by Jean-Antoine Nollet. For the following 200 years, osmosis was only a phenomenon observed in the laboratory. In 1949, the University of California at Los Angeles first investigated desalination of seawater using semipermeable membranes. Researchers from both University of California at Los Angeles and the University of Florida successfully produced fresh water from seawater in the mid-1950s, but the flux was too low to be commercially viable[2] until the discovery at University of California at Los Angeles by Sidney Loeb[3] and Srinivasa Sourirajan at the National Research Council of Canada, Ottawa, of techniques for making asymmetric membranes characterized by an effectively thin "skin" layer supported atop a highly porous and much thicker substrate region of the membrane. By the end of 2001, about 15,200 desalination plants were in operation or in the planning stages worldwide.

Osmosis is a natural process. When two liquids with different concentrations of a solute are separated by a semipermeable membrane, the fluid has a tendency to move from low to high solute concentrations for chemical potential equilibrium.

Formally, reverse osmosis is the process of forcing a solvent from a region of high solute concentration through a semipermeable membrane to a region of low solute concentration by applying a pressure in excess of the osmotic pressure. The largest and most important application of reverse osmosis is the separation of pure water from seawater and brackish waters; seawater or brackish water is pressurized against one surface of the membrane, causing transport of salt-depleted water across the membrane and emergence of potable drinking water from the low-pressure side.

The membranes used for reverse osmosis have a dense layer in the polymer matrix—either the skin of an asymmetric membrane or an interfacially polymerized layer within a thin-film-composite membrane—where the separation occurs. In most cases, the membrane is designed to allow only water to pass through this dense layer, while preventing the passage of solutes (such as salt ions). This process requires that a high pressure be exerted on the high concentration side of the membrane, usually 2–17 bar (30–250 psi) for fresh and brackish water, and 40–82 bar (600–1200 psi) for seawater, which has around 27 bar (390 psi)[4] natural osmotic pressure that must be overcome. This process is best known for its use in desalination (removing the salt and other minerals from sea water to get fresh water), but since the early 1970s, it has also been used to purify fresh water for medical, industrial, and domestic applications.

2.Application and system technical characteristics ARIUM-6131

Description

The compact arium® RO 61316  laboratory water purification system is designed to produce RO-Water for general laboratory applications.

With production volumes up to 16 l/h and up to 99% rejection rate of ions, bacteria, organics and viruses, automated RO-membrane back-flushing, and a constant flow rate, the arium® 61316 is the ideal choice for daily laboratory applications.

The efficient RO-membranes reduce water wastage and provide excellent product water quality, whilst ensuring high retention rates of the water impurities.

By the use of an optional closed pressurized tank (30, 70, 100 liter) the RO product water will be distributed with up to 3 bar to the point of use. This unique feature provides a pressurized product-water flow for all your lab applications. There is no need for an additional distribution pump.

Applications

- Feed water for ultrapure water systems

- Rinsing glassware and laboratory vessels

- Feed water for humidifiers, autoclaves and dish washer

Features

- Four line alphanumerical display

- Simple menu navigation

- Displaying conductivity, rejection rate, tank-level, temperature etc.

- Typical conductivity < 20 ^S/cm

- Production capacity up to 380 liter per day

- Easy to change pre-treatment cartridge and RO-modules

- Automatic RO-membrane-back-flush with RO-permeate

- Constant flow rate

- Several alarm functions

- Product water storage in open gravity or closed pressurized tank

- Serial interface RS-232

- PLC interface for external communication

3. Inlet Feed Water Requirements

Potable tap water feed only

| Minimum inlet pressure: | 1.2 bar¦ 15 psi |
| --- | --- |
| Maximum inlet pressure: | 6.8 bar¦ 100 psi |
| Temperature: | 5°C-30°C |
|  | (41°F-86°F) |
| Hardness (max CaCO3): | 360 ppm or |
|  | softened water |
| Silt density index: | < 3 |
| Free chlorine: | 0.1 ppm |
| Turbidity: | < 1 NTU |
| Iron (total as Fe): | < 0.1 ppm |
| Ambient Temperature & Humidity |  |
| Operation: | 5°C-30°C |
|  | (41°F-86°F) |
|  | 80% relative |
|  | humidity |
| Storage: | 5°C-45°C |
|  | (41°F-113°F) |
|  | 80% relative |
|  | humidity |

4. RO System Schematic

5.Operation of the reverse osmosis ARIUM- 61316

1. Device at all times must be connected to the network.

2 . Water should always be open. Outside weekdays water can be closed.

3 . Softener twice a week independently regenerated. At this time, the TOC may fail due to the ingress of reclaim system . This phenomenon of self-with a few minutes.

4 . Periodically monitor the level of salt in the softener . minimum

the amount of salt corresponds to 1/ 3 of the height of the tank.

5 . Regeneration TOC held 1 time in 3 months. If necessary, it will be displayed

message «PERFORM SANITISATION». TOC is used for regeneration of the syringe 612CDS.

6. Regeneration tank collecting purified water. Here we use a ready solution of

syringe 611 CDS /.

7. Average value of the conductivity of purified ( deionized ) water 20-60 S/cm. Siemens on average S/cm.

6.Procedure for switching on and off of the reverse osmosis

The order of the work :

1.Proverit water. It must be turned on ( open).

2 . On the display, press the button «OPERATE» / If the squeak will be issued , you must press the «ENTER».

To turn off

1. Click on «STANDBY». Wait for system shutdown . This will be

indicate silence and the appearance of the instrument display reads «STANDBY». TOC disabled. On Saturday , you can turn off the water .

Notes:

1. When the water is turned off or in a mode «STANDBY» approximately

every 3 hours independently activated for 15 minutes for washing. Thus when

a closed water tap water installation for flushing of the tank takes ( the collection of deionized water).

2 . In case of prolonged absence of laboratory staff ( more than 2 days ), you must

turn off the power after performing " Shutdown ."

3 . If the water pressure at the inlet to the plant will be small for inclusion in the installation work, then , in the absence of other causes, such as low pressure water supply in the water , replace the strainer . Height 10 inch filters or 250 mm. , Pore size 5 microns. To open the filter for spare parts has a special key with blue pen.

4 . Do not allow air into the system . If there is air in the water , then

it is necessary to release the spigot sink . If the air is turned in the system , it is

Modern exits the system as the installation work .

5 . Conductivity sensors are installed on the rear housing part. Measuring cells embedded in the original tube and purified water.

Requirements for qualification

Allowed to perform the study researches engineer , having higher education , experience in the chemical laboratory and read the instructions for appliances.

Safety regulations

When working with multiskanerom PCD- 650 spectrometer Aanalyst 800 prefix injection FIAS 100 and installing reverse osmosis ARIUM- 61316 must fulfill the general " rules of safety when working in a chemistry lab " and guided by regulations:

GOST - 12.1.0.019 Electrical while dealing with electric ;

- space laboratory must comply with the fire safety requirements in accordance with GOST 12.1.004 and have fire fighting equipment ;

- ventilation GOST 12.4.021-75 .

List of references

1.Integrated instruction manual WP 600 series meters. www. eutechinst.com

User Guide mnogoparametrnogo kiberskanera PCD 650.

2 . Alemasova A. Analytical Atomic Absorption Spectroscopy / A.

S. Alemasova , AN Rokun , IA Shevchuk . Sevastopol: Weber 2003 .

3.Pupyshev AA Practical Course atomic absorption analysis . course

lectures / AA Pupyshev . Yekaterinburg: HPE USTU , 2003.4 .

4 . Installation and Operation Manual ARIUM- 61316 http://www.vivaproducts.com/ Manual reverse osmosis ARIUM- 61316 .

Normative references

1. ГОСТ 6709-97 " distilled water ", water type III by ASTM, NCCLS, ISO 3696 , ATS; FS 42-2619-97 " Purified Water "

2 . ГОСТ 12.0.004-90 Occupational safety standards system .

Organization of training safety. General Provisions .

3 . ГОСТ 12.1.004-91 Occupational safety standards system .

Fire safety . General requirements .

4.ГОСТ 12.1.005-88 Occupational safety standards system . Common

hygienic requirements to the working area.

5.ГОСТ 12.1.007-76 Occupational safety standards system . harmful

substance. Classification and general safety requirements .

6.ГОСТ 12.1.019-79 Occupational safety standards system .

Electrical . General requirements and nomenclature of species protection.

7.ГОСТ 12.2.007.0-75 safety standards system . ware

electrical . General requirements for safety .

8.ГОСТ 12.4.009-83 Occupational safety standards system .

Fire fighting equipment to protect objects. Basic types. Accommodation and service

9.ГОСТ12.4.021 -75 safety standards system . systems

ventilation . General requirements .

10.ГОСТ 12026-76 Laboratory filter paper . Specifications.

11.ГОСТ 1770-87 Measured glassware Laboratory glassware . Cylinders , beakers, flasks, test tubes . Specifications.

12.ГОСТ 4204-77 sulfuric acid . specifications

Reagents 13.ГОСТ 4212-76 . Preparation of solutions for the colorimetric and nephelometric analysis.

14.ГОСТ 6709-72 distilled water . Specifications.

15.ГОСТ 7328-82 Measures mass general and exemplary . Specifications.

16.ГОСТ 18300-87 Rectified ethyl alcohol for industrial use. specifications

17.ГОСТ 21400-75 Glass chemical-laboratory . Technical requirements . test methods

18.ГОСТ 24104-2001 Laboratory balance of general purpose and exemplary . General specifications

19.ГОСТ 25336-82 glassware and laboratory glass . Types, basic parameters and dimensions

20.ГОСТ 29169-91 Laboratory glassware . One-mark pipettes (ISO 648-77 )

21 ГОСТ 20292 Laboratory glassware . Pipettes . (ISO 835/1 -81 ) Part 1 . general requirements

Besterekov U., Goncharov V.

COLLECTION OF METHODICAL INSTRUCTIONS

"DETERMINATION OF                                                                                 QUANTITATIVE AND QUALITATIVE INDICATORS OF WATER"

Signed for print ________________

Demy XxY  1/16

Typographical paper. Offset printing. Volume 55 p 3,4375 п.л.

Edition 20 copies. Order № ... *

Publishing Center SKSU. M.Auezov, Shymkent, Tauke Khan 5